GABA : Analytical Parameter
Parameter Value Additional information / Instrument Name Bruker Avance III / Magnetic field strength 500Mhz / Software TopSpin 3.2 / NMR probe BBi / 1D Acquisition proton EXPNO 1-2-3-4 / Pulse sequence ZG / Pulse Angle 90 / Number of scans 64 / Temperature 300K / D1 2-10-20-25s / 1D Acquisition proton EXPNO 5 / Pulse sequence Noesy1D / Pulse Angle 90 / Number of scans 16 / Temperature 300K / D1 2 / Mixing time 10 ms / Expected result Water suppression / 1D Acquisition 13C {1H} EXPNO 8 / Pulse sequence zgpg / Pulse Angle 90 / Number of scans 10240 / Temperature 300 / D1 2 / Decoupled garp / 2D Acquisition proton Pulse sequence jresgpprqf EXPNO 6 / Pulse Angle 90-180 / Number of points 8192 (F2) * 32 (F1) / Number of scans 8 / Temperature 300 / D1 4 / Expected result Multiplicity of protons and coupling constant J (Hz) / 2D Acquisition proton Pulse sequence cosygpqf EXPNO 7 / Pulse Angle 90 / Number of points 4096 (F2) * 512 (F1) / Number of scans 20 / Temperature 300 / D1 2 / Expected result coupling proton-proton / 2D Acquisition proton/carbone Pulse sequence hsqcetgp EXPNO 9 / Pulse Angle 90-180 / Number of points 4096 (F2) * 256 (F1) / Number of scans 48 / Temperature 300 / D1 2 / Expected result coupling proton-carbone / NMR tube : Preparation Method / Solute concentration mM 20.0 +/- 0.5 / pH If aqueous medium 6 / Solvent D2O / Buffer phosphate 200mM / Reference TSP 1mM / Data Processing 1D proton spectra and carbon are phased manually. The baselines are corrected automatically. The spectra were referenced to the TSP peak of which is set to zero automatically and manually for the proton on carbon. / Peak Picking The PP is initially performed automatically in the proton spectrum (expno 4) and on the carbon spectrum (expno 8). The list of peaks is obtained verified. If necessary, peaks which are not associated with the studied molecule (dirt, noise, etc ...) are removed manually. The data is exported directly from TOPSPIN3.2 to an EXCEL file. / Minimum relative Intensity 0.0001 / Maximum relative Intensity 10000 / Peak detection sensitivity 1 / PPM range 11 … -0.1 / Intensity of the reference peak for the PP(most intense peak) 15 / Peak List Table of chemical shifts and intensities with attribution Excel File GABA.xls / Peak tab Proton spectrum peak list / Peak+sat tab Proton spectrum peak list + 13C satellite detected / Multiplet tab Assignment of protons of the molecule. The coupling constants J in red mean that data values ​​were measured directly on the JREs. / Carbone tab Carbon spectrum peak list. The annotation column gives the assignment to which the HSQC was used / HSQC tab Coordinates spots correlation